## Slide 1: Noninvasive Glucometer
General Updates
Nov 2024
https://github.com/knowthebird/NoninvasiveGlucometer

## Slide 2: Contest Updates
There were no contest entries 
  Prizes will be saved/used as fund for noninvasive glucometer projects
  If you have an idea for a design and want to chat about it, or request some support, or if you are interested in refining the designs in the git repo, reach out to noninvasiveglucometer@gmail.com
There were people who reached out with interest in how things were designed and suggestions for improvements in the future 
Just because there were not entries in the contest does not mean there is not interest in the problem
  Many are going to look to protect their intellectual property, and not quickly share it for collaboration

## Slide 3: New Log Data
Lots of new samples,
  89, two logs per session, separate fingers
  A very wide range of concentration, 76 – 447 mg/dL
  Thank you Dave!
Expected results are:
  As blood glucose concentration increases more light is absorbed by the glucose and both reflectance and transmittance sensors report less light
  For this design, sensors report higher voltage when more light is absorbed
    12-bit resolution, 3V ~= 4096, 0V = 0
  When looking at results, higher number should mean higher concentration
This was not observed.
  Even when looking at the same finger and same sensor, the general voltage does not follow the glucose concentration.

## Slide 4: Glucosan 1- Hardware Failure
PPG Waveform appears to be
Duplicated in readings.

## Slide 5: Glucosan 1- Hardware Failure
Signal appears to become duplicated .
Occurs in state where LEDs or all off.
Duplication appears to stop.
Occurs in state where one LEDs is on.
Duplication appears to be present in
almost every log for 85-174. In some log
files it appeared to “toggle” off and on.
Note, there are not new software calls that
should be made during this “toggle”.

## Slide 6: Glucosan 1- Hardware Failure
Duplication becomes less seperated.
Occurs in state where LEDs or all off.
Duplication appears to widen again.
Occurs in state where two LEDs are on.
Even when less separated, there does appear
to be duplication or shifting occurring here.

## Slide 7: Glucosan 1- Hardware Failure
In some cases, the sensors appear to be completely failing.

## Slide 8: Movement Artifacts
Easy to see visually looking at logs but will require post processing algorithms to filter out.
Currently no IMU to monitor for movement for post processing or to alert.

## Slide 9: In Hindsight – Lesson Learned
Getting a variety of samples is not trivial
  Glucose concentrations typically stays in a set range, and extremes are a health risk.
Hardware needs to be more reliable before sending out in the world
A way of sharing results remotely (ex: over wifi), could make it easier to remotely identify issues faster

## Slide 10: Short Term Next Steps – Simplify the Problem
Measure Controlled Samples In A Controlled Environment
  Cuvette containing solution of dextrose monohydrate in distilled water
  Fixed position of emitters and sensors
    Can 3D print a jig
  PC Oscilloscope for Sampling Sensor
    Built in Logging
      No concerns of microcontroller wiring or logging corrupting sampling
    High Sampling Rates
      Will be more helpful when ready to monitor PPG waveforms
  Test with probes and consider coverings to reduce noise in readings.

## Slide 11: Longer Term – Improvements For A Portable Device
Fixed position for Sensor and Emitter
  Makes problem less complex when not accounting for their movement
IMUs on device and/or finger/ear
  Need to be able to say what are artifacts of movement and be able to notify user of bad samples
Maybe Cover the sensor and emitter
  Maybe you a bag, glove, or cover of some kind to prevent noise on sensors
  May be more important with a fixed sensor/emitter introducing air gaps
Very long term, consider ability to share data through wifi or cell phone application
  Could allow quicker results, and possibly a method for pushing updates back to the device
  Adds complexity and data security risks too

## Slide 12: The Research and Development Is Active
Companies
  Abbott Laboratories
  Apollon
  Ascensia Diabetes Care Holdings AG
  Biolinq Inc.
  DexCom, Inc.
  DiaMonTech AG
  Glucosense Diagnostics
  HagarTech
  Integrity Applications, Inc.
  Know Labs
  LifePlusAI
  LifeScan, Inc.
  Medtronic plc
  Movano, Inc.
  Nemaura Medical
  Profusa, Inc.
  Roche Diagnostics
  Scanbo
  Senseonics Holdings, Inc.
  RSP Systems
  PK Vitality
  Cnoga
Universities
  Cardiff University
  Kennesaw State University
  Massachusetts Institute of Technology (MIT)
  Monash University
  National Taiwan Ocean University
  RMIT University
  Singapore National Population Health Survey
  University of Chinese Academy of Sciences
  University of Leeds
  University of Missouri
  Ulsan National Institute of Science and Technology (UNIST)
Lists provided by Amer Khaqan